JJQ-L2016

Semi-automatic Cryostat Microtome

Instructions manual

Attention: Before using the machine, please read this Instruction carefully,

and keep the Instruction for further reference.

WUHAN JUNJIE ELECTRONICS CO., LTD

Address：Chuangye Street 10-A-404, Guanggu, East Lake High-tech Development Zone, Wuhan City

Postcode：430074

Tel: 8627-87401756 / 862787411993        Fax:8627-87411993

Website：http://www.whjjbl.com		   E-mail： junjiedianzi@126.com

Content

1.Introduction	 3

- Scope of application	3

- Consists of parts	 3

- Helpful suggestions	3

2.Safety Instructions	 3

- Obligation to the operator	 3

- Safety notes	 4

- Technical Parameter	 4

  - Technical Specification 	 4

  - Feature	 4

  - Dimension	5

- Working Condition	 5

  - Advanced set	5

  - Ambient temperature and environment	 6

- Assembling and parts adjustment	 6

  - Specimen head	 6

  - Blade holder	 7

    - Replace blade	8

    - Adjust the angle of the blade	 8

  - Adjust the anti rolling plate	 8

    - Make the edge of the anti rolling plate parallel to the edge of the blade	 9

    - Adjust the anti rolling plate angle with the blade	 9

    - Make the edge of plate and edge of blade meet flush	10

  - Hand wheel locking	10

  - Cleaner	11

6.Operation	 11

- Turning the Instrument on	 11

- Operation panel	 11

- Left area control buttons introduction	 12

- Middle area control buttons introduction	 12

- Right area control buttons introduction	 14

- Usage of small operation panel	 16

- Trouble shooting	 18

- Safety and maintenance	 19

- Cleaning	19

  - Cleaning up the instruments	19

    - Conduct the following steps before cleaning each time	 19

    - Instruments and external surface	 20

    - Blade holder	20

  - Lubricate instruments	 20

    - Instruments and specimen holder	 21

    - Blade holder	21

Notes ：Slice Reference	 21

1.Introduction

Model JJQ-L2016 Semi-automatic Cryostat Microtome, national regulated 1st class medical equipment, is a rapid pathological section instrument to human and animal body tissues. It can be widely used for pathological diagnosis, analysis and research in hospitals, medical colleges, legal medical experts and propagation institutes.

- Scope of application

Used to Pathological section and study histology of plant, animal or human being.

- Consists of parts

The instrument mainly consists of 4 parts:

- The computer controlled panel is on the upper part to display the temperature and the working condition of the instrument.

- The center part is the constant low temperature room to rapidly freeze the vivid tissue and to conduct the cutting area.

- Behind is the mechanical transmission and motor drive.

- The lower part is the compressor unit for cooling.

- Helpful suggestions

- The user should read this user manual carefully before first installation and operating the machine, and put this manual beside the instrument to refer it when meet problems.

- Before operation, please carefully read operation instruction first to know operation steps and process.

2.Safety Instructions

- Obligation to the operator

Only staff who are familiar with basic regulations on safety at work and accident prevention and are trained in the use and operation of this instrument should be allowed to handle this instrument.

- Safety notes

- Unit must ship upright. If this unit was shipped lying down or stored lying down then it must stand upright for 36 hours before use.

- Do not put instrument under extreme temperature and high air humidity environment. Failure to follow this will cause instrument severe damage

- Blade is sharp, beware when you change it.

- Do not place blade anywhere with the cutting edge facing upwards.

- Check your power socket is well grounded. Ensure that power supply is stable, constant and adequate. Compressor requires a start-up current between45 and 50A. Therefore, the professional electrical engineer must inspect the electric circuit before installation to meet smooth operation of the instrument requirement.

- Do not place instrument directly under sunshine or outlets of air condition

- No other obstruct around the instrument in 30cm distance.

- Before changing specimens always lock the hand wheel .

- In case of malfunction, contact our company. Don't try to solve it by your own risk.

- Technical Parameter

3.1 Technical Specification

- Voltage：AC 220V±22V

- Frequency：50Hz±1Hz

- Power：600W

- Refrigerant：R404a,300±10g

- Compressor Type：Double Danfoss Compressor

- Working noise: 52dB(A)

- Feature

• Section thickness range： 1-100pmAdjustable

1-20pm , Increment： 1pm 20-40pm, Increment： 2pm 40-100pm， Increment： 5pm

- Trimming thickness range：	10-400pm Adjustable

10-50pm , Increment： 5pm 50-100pm , Increment： 10pm 100-400pm ， Increment： 50pm

- Specimen retraction： 0-80pm Adjustable ， Increment： 5pm

- Maximum specimen size： 45mmx50mm

- Horizontal stroke： 20mm Vertical stroke： 60mm

- Chamber temperature：	-10°C〜-40°C Adjustable

- Specimen head temperature：	-10C〜-40C Adjustable

- Freezing shelf ： Minimum Freezing shelf temperature：	-45C

Number of freezing station：	26

- Peltier： Minimum Peltier shelf temperature： -60C

Peltier number：	2

Working time to reach -60C： 15min

- Dimension

- Size： 660mm*640mm*1150mm	135kgs

- Packing size： 790mm*810mm*1336mm	175GK

- Working Condition

  - Advanced set

• The instrument is movable floor type with 4 caster wheels. Two front wheels can be up and down adjusted by turning knobs on the foot. These front wheels only function when the instrument needs to move. After remove the package, put the front wheels on the ground and move it to the working position and secure it, then disable the front wheels.

- The instrument can only be used after 4 hours standing still.

- Confirm the power supply is grounded (triple socket), then connect the instrument with the main power socket, turn on the power switch at the back of the instrument, the machine is in operation mode.

- Other materials are not allowed to put on both sides of the instrument. At least to leave a 300mm space to facilitate ventilation and heat dispersion.

  - Ambient temperature and environment

- The instrument working ambient temperature is +5oC〜+28oC.If the temperature is above 26OC, air conditioning must be installed.otherwise the life time of the cooling compressor will be shortened.

- Environment relative humidity is no more than 80%.

High temperature or humidity will influence freezing effect.

- Assembling and parts adjustment

  - Specimen head

Vertical adjust

Horizontal adjust

Fixed specimen

-10C~-40 C adjust

- Blade holder

Clamping lever for moving blade holder forward and backward

Handle for replacing

slide

Fastening slide knob

Clamping lever for locking blade

Turning lever to move blade holder

knob

left and right

Turn rotator to adjust

cutting angle of blade

• Release(anti clockwise)clamping lever on the left，manually move the holder back and lock it at the desired position. When the holder moves to the terminal point the whole blade holder can be taken out and put it back by reverse procedures. Release(anti clockwise)lever on the right, manually move the holder left and right and lock it at the desired position.

Note: Remember to lock well the blade holder to right position when doing forward or backward specimen adjustment.

- Replace blade

• Releaseclampingleverontheupperrighttoopenbladeclearance;insertbladefromlefttoright; then lock well the clamp lever.

Note：Blade is sharp, beware when you change it.

- Adjust the angle of the blade

• Use hexagon spanner to loose the rotator. Knife folder can be turned and in the selected location can be tight. Slice half-way when necessary change blade, only need to release clamping lever for locking blade, then replace the blade.

5.3Adjust the anti rolling plate

Anti rolling plate height turning knob

Using hexagonal wrench to make

Paraller bwtween blade and slide

Parallel

Blade

Knob for clamping anti rolling plate

The anti rolling plate is made of Plexiglas. Adjustment includes：

- Make the edge of the anti rolling plate parallel to the edge of the blade

- Insert blade and lock the clamping lever

- Lock two knobs.

- Using hexagonal wrench to make parallel between blade and anti rolling plate

- Fastening hexagonal screw

- Turing height turning knob to adjust height of anti rolling plate

5.3.2Adjust the anti rolling plate angle with the blade

- Make the edge of plate and edge of blade meet flush

Anti rolling plate height turning knob

• To protect the upper part of the anti-rolling plate connecting with tissue specimen without any damage.

• Whenthequalityofthetissuesectionisnotgood,Shouldviewthecuttingedge,anti-coilwhether organized debris, grease and other foreign matter.

• Avoidtotouchingthefrontoftheplate,becausehightemperaturewillstickthetissuesection.

- Hand wheel locking

Rotary lever

Hand wheel locker

Locking spanner

To give operator more convenience and safety, we increase locking position from 8 to 16. The first locking position is at the top while gearing handle is facing to operator.

- Cleaner

To Keep The Freezing Chamber And Operational Environment Clean And Tidy.

6.Operation

- Turning the Instrument on

First connect power line on the rear of cryostat, then turn on switch on the rear of cryostat. Cryostat is under working condition and feeding arm go back original position. Compressor starts to work after turning on power switch and freezing chamber begins to cool down. (It depend the interval of turning off and on, if within 3 minutes, the instrument needs to wait another 3 minutes, otherwise machine easy to damage). Two red LED lights above right area , first one is for main compressor and second is small compress. When compressor starts to work, the first light is on. If press fast freezing button, second light is on.

- Operation panel

Control board is separated into three area, left area, middle area and right area. Left area is for software switch and fast defrost buttons; middle area is for section/trimming setting, peltier and stand by function control; right area is for time and temperature setting. Small control panel on the left of cryostat controls clamp feeding, illumination and UV sterilizing.

Left area

Middle area

Right area

- Left area control buttons introduction

- Software switch

Pressthisbuttontoswitchonoroffoperationpanel.Ifindicatorison,usermay adjust time, defrost, temperature and so on. Otherwise, indicatoris offand panel

locked. User cannot do any operation.

- Immediate defrost

This is fast defrost button. Press it to activate fast defrost function and indicator light is on, system goes into fast defrost mode. It takes about 6 minutes to defrost. Press the button

again to stop fast defrost working, indicator light goes out simultaneously. Freezing recovers after 25 seconds. If system automatically enter time defrost mode, indicator lights up too. Press button again to stop defrost and indicator light goes out.

6.2.2Middle area control buttons introduction

6.2.2.1Peltier freezing ( It is invalid under stand by mode)

This is peltier switch button. Press it, indicator light is on and peltier element begins to work and cools down the stations. Forth line on the screen shows [015] to indicate peltier working 15 minutes. The countdown of remaining cooling time is displayed, such as [014], [013] 	”[002]”，“[001]”，until peltier stops

freezing and indicator off. During Peltier working time. User may press this button to stop peltier when it is activated. In addition, if freezing shelf temperature is higher than -10 °C when peltier working,peltier will self close. On the same principle, if freezing shelf temperature is higher than -10C, peltier cannot be activated.

- Standby

• We call the instrument into “Standby” mode by keeping the freezing chamber at a temperature of -4C to -9C . This time the operational time of the compressor is less than the stop time with average power consumption decreased, more prolongs the lifetime of the compressor. When long time do not

use the instrument for slicing, then let it go into “Standby” mode.

This is standby button. Press it one time, system goes into standby mode or transfer to work state. LCD screen background light tells us how to differentiate from standby and normal

working. If background light is in dark, it is standby mode, On the contrast, it is working.

- Adjust value

| Sect | 6 |
| --- | --- |
| Retr | 15 |
| No | 0 |
| Tn | 0 |

Picture 1

After switch on cryostat, the first LCD will display data as Picture 1. First line Sect is to represent section and its value; second line Retr is to represent retraction and its value; third line No is to represent total pieces and its value; forth line Tn is to represent total thickness and its value.

This is mode selection button. Press it may select Section (Sect) or Trimming(Trim) working mode.

“C” is clear button. Push it to clear total pieces and thickness into “0”.

• In addition, these two buttons are combination keys. If user push them together, there is one “*” signal appearing in front of Rect on the second line of LCD to show it under adjustable condition. User may press “+” or “-“ to change retraction value.

“+” button is to increase value. Push it to add on section (trimming) value. If “*” signal appearing in front of Rect, it is to adjust retraction value.

“-” button is to decrease value. Function is same as above.

6.2.3Right area control buttons introduction

- Clock adjust

| Day | 1 |
| --- | --- |
| Time | 12:00 |
| T0 | 16 [-18] |
| T1 | 16 [off] |

After cryostat switching on, it is self running under current time state, right LCD on control panel is showing information and data as picture 2. We call it initial mode when cryostat get power supply and reset. First line Day is to present current operational day, numbers from 1 to 7 separately stand for Monday to Sunday.

Second line Time is to present current operational time. Third line T0 is to present chamber temperature. It is actual temperature value for operational room; following value in [ ] is setting temperature of freezing chamber. Forth line T1 is freezing shelf temperature. It is actual temperature value for freezing shelf. “Off” in [] is to stand for peltier off. User can press “D” to adjust working day. Then press hour button ”H” and minute button “M” to adjust hour and minute respectively.

This button is to select temperature. Press button per time, the third line on the LCD will shift between T0( Chamber temperature) andT2 (Clamp temperature).

- Temperature setting

Press this button to increase temperature value. Value in [ ] in the third line onLCD will

add on. If LCD shows T0, it is to adjust freezing chamber setting; otherwise it shows T2 to adjust

clamp temperature.

This is decreasing button. Function is same as above.

- Fast freezing (It is invalid under standby mode)

• This cryostat has double compressors. Big compressor freezes for chamber and shelf. Small compressordoesnotworkifsystemisinstandby,whensystemgoesbacktonormalworkingfrom stand by, fast freezing is activated. Small compressor is going to start in 3 minutes and freeze clamp. Operator may cancel fast freezing regarding their own situation.

Press fast freezing button, button green indicator is on and with sound of compressor initializing. At same time, red LED on the panel for small compressor indicator should be on.Clamp temperature gradually decrease with working ofsmall compressor. Press this  buttonagain,twoindicatorsareoffandsmallcompressorstopworking.

- Programming start-of-work time

Time mode button  Working day button

Adjust day	Adjust hour  adjust minute

- If second line on the LCD is in Time mode,Press time mode button until second line show “TON”.Thenusercanpress“D”toadjustworkingday.Press“D”buttononetime,working day will minors one day. Then press hour button ”H” and minute button “M” to adjust hour and minute respectively. (Definite time interval is 10 minute, automatic defrost is same)

- User can press “W-DAY” button to set single day start-of-work time too. After push “W-DAY” button, first line in the LCD should change to “Dwork 1&6”. “1&6” represents Monday to Saturday. Press “D”button to select working day, then press ”H”and “M”buttons to adjust hour and minute respectively. System automatic set 24 hours for nonworking day and goes into standby mode to reduce electricity consumption.

- Programming end-of-work time

- Press time mode button as many time as “TOff” displaying on the second line on LCD;it stands for system enter end-of-work time programming; setting method is same as start-of-work programming. After operation press time mode button again to return working mode, and lock start/end setting time.

- Note: Please make LCD go back to initial condition mode after setting.

6.2.3.6 Programming defrost definite time

This is defrost definite time button. Under initial mode after pressing, first line on the LCD displays “Defrost”. Second line displays “Time”. User may press ”H” and “M” buttons to adjust hour and minute. If it is needless, please adjust hours to 24. Press button again to go back initial mode and lock defrost definite time.

6.2.4.1 UV sterilize

6.2.4 Usage of small operation panel

Thisis UV sterilize button. Press it to activate UV lamp and Indicator light is on, UV lamp in the freezing chamber is on and begins to sterilize. It is automatic timing after activating, and it will self close in 35 minutes. User may press button once more to stop it at any time.Indicator will be off correspondingly.

Note ： Keep quartz tube cleaning is the key point for sterilizing, We suggest user to regularly clean it to maintain sterilizing ability. Because UV sterilizing is unfit for low

temperature, It is better to start to sterilize under stand by situation.

Note: When doing UV sterilize, please close glass window and do not let UV light goes out of chamber.

- Lamp

This button is on and off button. First push it the lampis on; second push it the lampis off.

- Slow Feeding

Specimen position is adjustable by moving forward or backward. Press ‘丁’ button, specimen is moving to user; press “t” button, reverse moving.

6.2.4.4 Fast feeding

Press double arrow buttons to fast move clamp forward or backward. When user press backward button, clamp will go back to backward limitation, if need to stop on the midway, please lightly touch forward double arrow button, then clamp will stop moving Alarm will

warn while specimen is moving to front or end limitation.

▲

Note: When adjust specimen forward or backwards movements, the hand wheel must lock at the right position.

- Use of the wasted liquid barrel

• Use the wasted liquid barrel to collect defrost liquid, tissue thawing solution, flushing liquid and wasted objects. In order to prevent the human and environment hazard from the wastes, put 200ml 10% formalin or other sterile solution in the barrel in advance. The disposal of the wastes shall follow the related regulations of the hospital. Disconnect the tube that link with the freezing chamber, then take off the barrel.

- Trouble shooting

| Trouble | Solving Method/ Reason |
| --- | --- |
|  | Select thinner tissue, |
| Tissue crack | The freezing time is not quick enough Apply more embedding material |
| Tissue specimen off | The cutting tissue is soft The temperature of the specimen disc is low Ensurethefasteningoftheblade |
| Tissue moves on, but without slicing | Fix the specimen clamp Ensure the fastening of the specimen disc on the clamp Increase the blade angle |
| Tissue section rolling | Check the gap between the anti-rolling plate and the blade Increase the height ofthe anti-rolling plate Checktheangleoftheanti-rollingplateandtheblade |
| Tissue becomes soft when cutting | Prolong the refrigeration time of the anti-rolling plate and the blade |
| Tissue section stick to the anti-rolling | Use short hair brush to clean the anti-rolling plate |
| plate | Prolong the refrigeration time Clean the blade with hair brush |
| Tissue section overlapping | The clearance angle between the blade and the anti-rolling plate is too narrow |
| Vertical crack of tissue section | Edge of the blade has defection, need grinding Foreign matter on the blade |
| Tissue section narrowed | Adjust the clearance angle of anti-rolling plate and the blade evenly |
| Tissue section broken | Too low temperature of the tissue, too long time freezing, the tissue is affected |

| Vibration when slicing | Unfixed blade The angle of the blade is too big or too small The specimen disc is not fixed well Cutting speed is too fast |
| --- | --- |
|  | Dull blade |
| Uneven thickness of the tissue section | Angle is too small The specimen disc is not fixed well |

- Safety and maintenance

- Make hand wheel in the lock position when the instrument stops.

- Open the glass door after the instrument stops, to let the steam vaporized and keep the working room dry.

- Change the broken light accordingly (220 V 15W)， pay attention to safety when doing the exchange.

- The elements of the electric system, such as control panel, electric control box, only use expert for the maintenance, others no touch.

- Thecoolingsystemmainlyconsistsofcompressor,condenser,filter,andevaporator.The condenser (with fan) may accumulate dust after long time usage, which can affect the Cooling function. It is necessary to open the side ventilation board and use brush to clean the dust on the projecting vane ofcondenser, oruse high pressure airto blowit.

- Frequently clean the constant temperature chamber, to keep it clean. Pay special attention to not hurt by the blade.

- Cleaning

  - Cleaning up the instruments

    - Conduct the following steps before cleaning each time

- Turn up the specimen grip to the top and lock the hand wheel.

- Release the specimen grip and pull it out.

- Pick offthe blade from the blade holder and put it back to the blade box.

- Dismount the blade holder and its seat to clean up.

- Take down the specimen from the specimen nip. Clear away the section waste with dry brush.

- Take down the specimen grip to clear up separately.

    - Instruments and external surface

• If necessary, the external painted surface can be cleaned with light-duty commercial housework cleaner or suds. And then use wet cloth rub it until dry.You may use the substitute of xylene,

grease oil, grease scavenger to erase residual.The instruments must be dry when use again.

    - Blade holder

Please according to following steps to clean up the blade holder if it had been dismounted.

- Turn the eccentric rod handle in the lateral of the body of revolution and draw it out from sideward.

- Push the knife clamp back which have knife clip and shift it out from the rotary unit.

- Turn the eccentric rod handle in the lateral of the knife clap and draw it out from sideward.

- Dismount the knife clamp.

- Clean up all parts of the blade holder.

- Make the blade holder dry and assemble it together.

- Apply to thin layer of lubrication after clean up the parts which had been taken off.

- When fix the knife clip, make sure that its upper part is parallel with the back edge of the knife clamp seat.

▲

Don't use xylene or alcoholic liquid (e.g.: glass cleaner) when clean up.

  - Lubricate instruments

Do oil lubrication for the following parts monthly. (1~2 drop is well enough)

    - Instruments and specimen holder

- Grip draw in the clamp.

- Lock in iron at the “T” knife clap back of the microtome bed plate.

- The blade holder slide way on the microtome bed plate.

    - Blade holder

- Lock in iron at the “T” body of the rotary units on the knife clamp seat.

- The knife control grip is shift to the eccentric rod handle.

- The iron locking head on the knife clamp of the ”T” body of rotary unit and the knife holder with slide way.

- The grasping joystick of the blade.

Notes ：Slice Reference

Note 1 ：When use the cryostat microtome, better to control the blade speed and correctly adjust the anti-rolling plate is an important factor to cut out high quality sections. The proper speed of the blade can only be acquired by practice experience, and the skillful control of the hand wheel. The adjust position of the anti-rolling plate is more. Sometimes they interfere with each other; it needs careful adjustment without random putting them on the blade edge.

Note 2 ：When freeze the living tissue, the water content in it will condensed into ice, the tissue will be hard. The hardness of it changes with the temperature, the lower temperature changes the harder the tissue becomes. Different tissue under different temperature, to have high quality sections, this skill can only be got through practice. Most cutting temperature of the specimen tissue without fat and formalin, better to select between -13C〜-23C .

Note 3 ：In order to obtain high-quality pieces of tissue, pay attention to the following:

- Proper selection of working temperature in the chamber

- Correct slicing operation

- Fine adjustment of the anti-rolling plate

- Sharp slicing knife, correct cut angle

Note 4 ：The freezing section is just opposite to paraffin slicing，the freezing section is not cut the long edge, better to cut the shorter edge which means the line of contact is shorter.